## AirPressure
AirPressure Compensation and ConCoef1 calculation for Oxygen Optode
Insert actual AirPressure | 1013.25 | hPa | From local met office
Insert old ConCoef1 | 1.1059 |  | From sensor User Maintenance
Air Saturation from sensor | 98 | % | From Actual reading in Air
New Calculated ConCoef1 | 1.1284 |  | Save this value to sensor
Nominal AirSaturation | 100 | %

## Virtual AirPressure
2-point Calibration
Insert Airsaturation from Reference Sensor | 98.6 | % Air Saturation
Calculated Air PressureVirtual | 999.0645 | hPa
Set Passkey(1000)
Set Interval(30)
Save
Set Passkey(1000) |  | In bath with bubbling
Do CollectCalDataSat
Set Passkey(1000)
Set CalDataAPress (..) | 999.0645
Save
Set Passkey(1000) |  | In bath with NatriumSulfitt
Do CollectCalDataZero
Set Passkey(1000)
Do Calibrate

## CondCoef
Cell Factor Calibration
CellCoef Old | 4.456
Cond ref | 34.5
Cond Read | 34.3
CellCoef New | 4.482