## PCS(3)_wt485
WT485 Reference List (PCS(3))
Rivo I - III | [5 empty] | Ver.: V2.1.1
2026-05-21
Addr. | Bit | Type | No. of bytes | Access | Designation | Description
Cl2
0 |  | ASCII | 12 | R | Date of the software version | " V: A_08/95 "
1 |  | ASCII | 28 | R | Module name | "COMMON Electronic - Analyser"
2 |  | UINT16 | 2 | RW | Password of the interface | 0...999
3 |  | ASCII | 12 | R | Module type | PCS-Cl2
4 |  | UINT8 | 1 | R | Mode | not supported
5 |  | DS1 | 12 | R | Measured value Cl2 | Data structure 1
|  | INT16 | 2 |  | Value
|  | INT16 | 2 |  | Range start
|  | INT16 | 2 |  | Range end
|  | ASCII | 5 |  | Unit
|  | INT8 | 1 |  | Divider
6 |  | DS1 | 12 | R | - | not supported
|  | INT16 | 2 |  | Value | not supported
|  | INT16 | 2 |  | Range start | not supported
|  | INT16 | 2 |  | Range end | not supported
|  | ASCII | 5 |  | Unit | not supported
|  | INT8 | 1 |  | Divider | not supported
7 |  | DS1 | 12 | R | Measured value temperature | Data structure 1
|  | INT16 | 2 |  | Value
|  | INT16 | 2 |  | Range start
|  | INT16 | 2 |  | Range end
|  | ASCII | 5 |  | Unit
|  | INT8 | 1 |  | Divider
8 |  | DS1 | 12 | R | Limit 1 | Data structure 1
|  | INT16 | 2 |  | Value
|  | INT16 | 2 |  | Range start
|  | INT16 | 2 |  | Range end
|  | ASCII | 5 |  | Unit
|  | INT8 | 1 |  | Divider
9 |  | DS1 | 12 | R | Limit 2 | Data structure 1
|  | INT16 | 2 |  | Value
|  | INT16 | 2 |  | Range start
|  | INT16 | 2 |  | Range end
|  | ASCII | 5 |  | Unit
|  | INT8 | 1 |  | Divider
10 |  | DS1 | 12 | R | - | not supported
|  | INT16 | 2 |  | Value | not supported
|  | INT16 | 2 |  | Range start | not supported
|  | INT16 | 2 |  | Range end | not supported
|  | ASCII | 5 |  | Unit | not supported
|  | INT8 | 1 |  | Divider | not supported
11 |  | DS1 | 12 | R | - | not supported
|  | INT16 | 2 |  | Value | not supported
|  | INT16 | 2 |  | Range start | not supported
|  | INT16 | 2 |  | Range end | not supported
|  | ASCII | 5 |  | Unit | not supported
|  | INT8 | 1 |  | Divider | not supported
12 |  | DS1 | 12 | R | Setpoint | Data structure 1
|  | INT16 | 2 |  | Value
|  | INT16 | 2 |  | Range start
|  | INT16 | 2 |  | Range end
|  | ASCII | 5 |  | Unit
|  | INT8 | 1 |  | Divider
13 |  | UINT8 | 1 | R | Moduloption | not supported
14 |  | UINT8 | 1 | R | Limit- and digitalinput | not supported
15 |  | INT16 | 2 | R | Error status | not supported
16 |  | INT16 | 2 | R | Calibration intervall | not supported
17 |  | FLOAT | 4 | RWP | Act. dosing rate/control rate | [%]
18 |  | UINT8 | 1 | R | Language | not supported
19 |  | INT16 | 2 | RWP | - | not supported
20 |  | INT16 | 2 | RWP | Setpoint Wi | Setpoint, [user defined unit], factor: 0.01
21 |  | INT16 | 2 | RWP | XP | [%]
22 |  | UINT16 | 2 | RWP | Tn | [min], factor: 0.1
23 |  | UINT8 | 1 | RWP | Limit 1 definition | not supported
24 |  | UINT8 | 1 | RWP | Limit 1 function | not supported
25 |  | INT16 | 2 | RWP | Limit 1 value | not supported
26 |  | INT16 | 2 | RWP | Limit 1 hysteresis | not supported
27 |  | UINT16 | 2 | RWP | Limit 1 delay | not supported
28 |  | UINT8 | 1 | RWP | Limit 2 definition | not supported
29 |  | UINT8 | 1 | RWP | Limit 2 function | not supported
30 |  | INT16 | 2 | RWP | Limit 2 value | not supported
31 |  | INT16 | 2 | RWP | Limit 2 hysteresis | not supported
32 |  | UINT16 | 2 | RWP | Limit 2 delay | not supported
33 |  | UINT8 | 1 | RWP | Limit 3 definition | not supported
34 |  | UINT8 | 1 | RWP | Limit 3 function | not supported
35 |  | INT16 | 2 | RWP | Limit 3 value | not supported
36 |  | INT16 | 2 | RWP | Limit 3 hysteresis | not supported
37 |  | UINT16 | 2 | RWP | Limit 3 delay | not supported
38 |  | UINT8 | 1 | RWP | Limit 4 definition | not supported
39 |  | UINT8 | 1 | RWP | Limit 4 function | not supported
40 |  | INT16 | 2 | RWP | Limit 4 value | not supported
41 |  | INT16 | 2 | RWP | Limit 4 hysteresis | not supported
42 |  | UINT16 | 2 | RWP | Limit 4 delay | not supported
43 |  | ASCII | 6 | RWP | Date / Clock | not supported
pH
0 |  | ASCII | 12 | R | Date of the software version | " V: A_08/95 "
1 |  | ASCII | 28 | R | Module name | "COMMON Electronic - Analyser"
2 |  | UINT16 | 2 | RW | Password of the interface | 0...999
3 |  | ASCII | 12 | R | Module type | PCS-pH
4 |  | UINT8 | 1 | R | Mode | not supported
5 |  | DS1 | 12 | R | Measured value pH | Data structure 1
|  | INT16 | 2 |  | Value
|  | INT16 | 2 |  | Range start
|  | INT16 | 2 |  | Range end
|  | ASCII | 5 |  | Unit
|  | INT8 | 1 |  | Divider
6 |  | DS1 | 12 | R | - | not supported
|  | INT16 | 2 |  | Value | not supported
|  | INT16 | 2 |  | Range start | not supported
|  | INT16 | 2 |  | Range end | not supported
|  | ASCII | 5 |  | Unit | not supported
|  | INT8 | 1 |  | Divider | not supported
7 |  | DS1 | 12 | R | Measured value temperature | Data structure 1
|  | INT16 | 2 |  | Value
|  | INT16 | 2 |  | Range start
|  | INT16 | 2 |  | Range end
|  | ASCII | 5 |  | Unit
|  | INT8 | 1 |  | Divider
8 |  | DS1 | 12 | R | Limit 1 | Data structure 1
|  | INT16 | 2 |  | Value
|  | INT16 | 2 |  | Range start
|  | INT16 | 2 |  | Range end
|  | ASCII | 5 |  | Unit
|  | INT8 | 1 |  | Divider
9 |  | DS1 | 12 | R | Limit 2 | Data structure 1
|  | INT16 | 2 |  | Value
|  | INT16 | 2 |  | Range start
|  | INT16 | 2 |  | Range end
|  | ASCII | 5 |  | Unit
|  | INT8 | 1 |  | Divider
10 |  | DS1 | 12 | R | - | not supported
|  | INT16 | 2 |  | Value | not supported
|  | INT16 | 2 |  | Range start | not supported
|  | INT16 | 2 |  | Range end | not supported
|  | ASCII | 5 |  | Unit | not supported
|  | INT8 | 1 |  | Divider | not supported
11 |  | DS1 | 12 | R | - | not supported
|  | INT16 | 2 |  | Value | not supported
|  | INT16 | 2 |  | Range start | not supported
|  | INT16 | 2 |  | Range end | not supported
|  | ASCII | 5 |  | Unit | not supported
|  | INT8 | 1 |  | Divider | not supported
12 |  | DS1 | 12 | R | Setpoint | Data structure 1
|  | INT16 | 2 |  | Value
|  | INT16 | 2 |  | Range start
|  | INT16 | 2 |  | Range end
|  | ASCII | 5 |  | Unit
|  | INT8 | 1 |  | Divider
13 |  | UINT8 | 1 | R | Moduloption | not supported
14 |  | UINT8 | 1 | R | Limit- and digitalinput | not supported
15 |  | INT16 | 2 | R | Error status | not supported
16 |  | INT16 | 2 | R | Calibration intervall | not supported
17 |  | FLOAT | 4 | RWP | Act. dosing rate/control rate | [%]
18 |  | UINT8 | 1 | R | Language | not supported
19 |  | INT16 | 2 | RWP | - | not supported
20 |  | INT16 | 2 | RWP | Setpoint Wi | Setpoint, [user defined unit], factor: 0.01
21 |  | INT16 | 2 | RWP | XP | [%]
22 |  | UINT16 | 2 | RWP | Tn | [min], factor: 0.1
23 |  | UINT8 | 1 | RWP | Limit 1 definition | not supported
24 |  | UINT8 | 1 | RWP | Limit 1 function | not supported
25 |  | INT16 | 2 | RWP | Limit 1 value | not supported
26 |  | INT16 | 2 | RWP | Limit 1 hysteresis | not supported
27 |  | UINT16 | 2 | RWP | Limit 1 delay | not supported
28 |  | UINT8 | 1 | RWP | Limit 2 definition | not supported
29 |  | UINT8 | 1 | RWP | Limit 2 function | not supported
30 |  | INT16 | 2 | RWP | Limit 2 value | not supported
31 |  | INT16 | 2 | RWP | Limit 2 hysteresis | not supported
32 |  | UINT16 | 2 | RWP | Limit 2 delay | not supported
33 |  | UINT8 | 1 | RWP | Limit 3 definition | not supported
34 |  | UINT8 | 1 | RWP | Limit 3 function | not supported
35 |  | INT16 | 2 | RWP | Limit 3 value | not supported
36 |  | INT16 | 2 | RWP | Limit 3 hysteresis | not supported
37 |  | UINT16 | 2 | RWP | Limit 3 delay | not supported
38 |  | UINT8 | 1 | RWP | Limit 4 definition | not supported
39 |  | UINT8 | 1 | RWP | Limit 4 function | not supported
40 |  | INT16 | 2 | RWP | Limit 4 value | not supported
41 |  | INT16 | 2 | RWP | Limit 4 hysteresis | not supported
42 |  | UINT16 | 2 | RWP | Limit 4 delay | not supported
43 |  | ASCII | 6 | RWP | Date / Clock | not supported
ORP
0 |  | ASCII | 12 | R | Date of the software version | " V: A_08/95 "
1 |  | ASCII | 28 | R | Module name | "COMMON Electronic - Analyser"
2 |  | UINT16 | 2 | RW | Password of the interface | 0...999
3 |  | ASCII | 12 | R | Module type | PCS-Redox
4 |  | UINT8 | 1 | R | Mode | not supported
5 |  | DS1 | 12 | R | Measured value ORP | Data structure 1
|  | INT16 | 2 |  | Value
|  | INT16 | 2 |  | Range start
|  | INT16 | 2 |  | Range end
|  | ASCII | 5 |  | Unit
|  | INT8 | 1 |  | Divider
6 |  | DS1 | 12 | R | - | not supported
|  | INT16 | 2 |  | Value
|  | INT16 | 2 |  | Range start
|  | INT16 | 2 |  | Range end
|  | ASCII | 5 |  | Unit
|  | INT8 | 1 |  | Divider
7 |  | DS1 | 12 | R | - | not supported
|  | INT16 | 2 |  | Value | not supported
|  | INT16 | 2 |  | Range start | not supported
|  | INT16 | 2 |  | Range end | not supported
|  | ASCII | 5 |  | Unit | not supported
|  | INT8 | 1 |  | Divider | not supported
8 |  | DS1 | 12 | R | Limit 1 | Data structure 1
|  | INT16 | 2 |  | Value
|  | INT16 | 2 |  | Range start
|  | INT16 | 2 |  | Range end
|  | ASCII | 5 |  | Unit
|  | INT8 | 1 |  | Divider
9 |  | DS1 | 12 | R | Limit 2 | Data structure 1
|  | INT16 | 2 |  | Value
|  | INT16 | 2 |  | Range start
|  | INT16 | 2 |  | Range end
|  | ASCII | 5 |  | Unit
|  | INT8 | 1 |  | Divider
10 |  | DS1 | 12 | R | - | not supported
|  | INT16 | 2 |  | Value | not supported
|  | INT16 | 2 |  | Range start | not supported
|  | INT16 | 2 |  | Range end | not supported
|  | ASCII | 5 |  | Unit | not supported
|  | INT8 | 1 |  | Divider | not supported
11 |  | DS1 | 12 | R | - | not supported
|  | INT16 | 2 |  | Value | not supported
|  | INT16 | 2 |  | Range start | not supported
|  | INT16 | 2 |  | Range end | not supported
|  | ASCII | 5 |  | Unit | not supported
|  | INT8 | 1 |  | Divider | not supported
12 |  | DS1 | 12 | R | - | not supported
|  | INT16 | 2 |  | Value | not supported
|  | INT16 | 2 |  | Range start | not supported
|  | INT16 | 2 |  | Range end | not supported
|  | ASCII | 5 |  | Unit | not supported
|  | INT8 | 1 |  | Divider | not supported
13 |  | UINT8 | 1 | R | Moduloption | not supported
14 |  | UINT8 | 1 | R | Limit- and digitalinput | not supported
15 |  | INT16 | 2 | R | Error status | not supported
16 |  | INT16 | 2 | R | Calibration intervall | not supported
17 |  | FLOAT | 4 | RWP | Act. dosing rate/control rate | not supported
18 |  | UINT8 | 1 | R | Language | not supported
19 |  | INT16 | 2 | RWP | - | not supported
20 |  | INT16 | 2 | RWP | Setpoint Wi | not supported
21 |  | INT16 | 2 | RWP | XP | not supported
22 |  | UINT16 | 2 | RWP | Tn | not supported
23 |  | UINT8 | 1 | RWP | Limit 1 definition | not supported
24 |  | UINT8 | 1 | RWP | Limit 1 function | not supported
25 |  | INT16 | 2 | RWP | Limit 1 value | not supported
26 |  | INT16 | 2 | RWP | Limit 1 hysteresis | not supported
27 |  | UINT16 | 2 | RWP | Limit 1 delay | not supported
28 |  | UINT8 | 1 | RWP | Limit 2 definition | not supported
29 |  | UINT8 | 1 | RWP | Limit 2 function | not supported
30 |  | INT16 | 2 | RWP | Limit 2 value | not supported
31 |  | INT16 | 2 | RWP | Limit 2 hysteresis | not supported
32 |  | UINT16 | 2 | RWP | Limit 2 delay | not supported
33 |  | UINT8 | 1 | RWP | Limit 3 definition | not supported
34 |  | UINT8 | 1 | RWP | Limit 3 function | not supported
35 |  | INT16 | 2 | RWP | Limit 3 value | not supported
36 |  | INT16 | 2 | RWP | Limit 3 hysteresis | not supported
37 |  | UINT16 | 2 | RWP | Limit 3 delay | not supported
38 |  | UINT8 | 1 | RWP | Limit 4 definition | not supported
39 |  | UINT8 | 1 | RWP | Limit 4 function | not supported
40 |  | INT16 | 2 | RWP | Limit 4 value | not supported
41 |  | INT16 | 2 | RWP | Limit 4 hysteresis | not supported
42 |  | UINT16 | 2 | RWP | Limit 4 delay | not supported
43 |  | ASCII | 6 | RWP | Date / Clock | not supported